## Slide 1: CSNE BCI Test Bed Design
Peter Xu
Yeongjee Chung

## Slide 2: Brain Computer Interface
Three Phases
  Phase 1 : External EEG and EMG measurements Off the shelf hardware, low power Bluetooth , 16 channels, Target board size 3 cm by 2 cm for 16 channel.
  Phase 2: (Implantable) Integrate the front end IA with the sensor array. Consider IR as an alternative communication mechanism to carry signals out of body.

## Slide 3
RHD2000 Features
Fully Integrated Electrophysiology front end chip
16-bit ADC
  30 kSamples/s/channel
32 or 16 channel input
  16 channel differential input
SPI interface with microcontroller
Low input-referred noise: 2.4 uV
DSP to filter out DC noise
Adjustable lower and upper cutoff frequency
Each amplifier on/off controllable
  Each amplifier consumes: 7.6uA/kHz
    Based on upper cutoff frequency
Electrical Characteristic
  VDD = 3.3v
  Baseline Amplifier current = 200 uA
  Amplifier current = 7.6uA/kHz
  Baseline ADC current = 510uA
  ADC current = 2.14uA/(kS/s)
  SPI current = 40 uA

## Slide 4: RHD2216 Block Diagram

## Slide 5: RHD2000 Characteristic
Notes: TOP : EMG signal recorded with RHD2216 using three Ag/AgCl electrodes during bicep contractions (in0+ and in0– on bicep, 5 cm apart; ground on elbow). Amplifier was configured with fL = 2.0 Hz, fH = 1.0 kHz, and DSP high-pass filter set to 10 Hz. ADC sampling rate was 4 kS/s per channel

## Slide 6: RHD2000 Interface with nRF51822

## Slide 7: nRF51822 Features
32-bit ARM Cortex M0 processor
256kB flash/ 16 kB RAM
31 GPIO pins
2 SPI channels
2.4 GHz radio
  Programmable output power +4dBm to -20dBm
  Tx and Rx radio can be enable separately
16 MHz oscillator
GFSK Modulation
2 Mbps maximum On-air data rate
-85dBm at 2 Mbps sensitivity
Selectable peripheral module On/OFF
  2 uA when all peripherals in IDLE and CPU OFF

## Slide 8
Electrical Characteristic (Processor)
  VDD = 3.0v
  I @ 16MHZ running off flash memory = 4.4 mA
Electrical Characteristic (Radio 2.4 ghz)
  VDD = 3.0v
  Tx Current @ +4dbm = 16 mA
  Tx Current @ -16dbm = 6 mA
  Rx Current @ 1 mbps = 13 mA
Electrical Characteristic (SPI)
  VDD = 3.0v
  Current @ 125kbps = 180 uA
  Current @ 8 Mbps = 200 uA
Mechanical Characteristic
  L x W x H = 6 x 6 x .85 (mm)
nRF51822 Characteristic

## Slide 9: Test Bed System
Block Diagram(Closed Loop Control)
Sensory Front-End
Amp
Pre-Filter
Post-Filter
I/O
A/D
Wireless RF
Wireless Optic
Actuator Front-End
External
Actuator
Robotics
Wireless RF
Wireless Optic
Signal Processing
Controller
Sensory/Actuator I/F
Control I/F
Wireless I/F
Bio-Signal I/F
I/O
D/A
Control Logic for Encoding

## Slide 11: Modular Building Block Structure
Feature Extraction
Prediction
Passive Motion Machine
Functional Stimulator
ZigBee
VCSEL
ECoG
EMG
EEG
Analog Process Module
A/D
BlueTooth
Haptic Feedback
DSP
Module
Wireless I/F
BioSignal I/F
Control Module
DSP Module
Comm.
Module
Comm.
Module
Comm.
Module
Brain/Body
Wearable BCI
BioData I/F
External Device
Multi-Ch.
Wrist Driven Orthosis
Prosthetic Devices

## Slide 12: Modular Software Building Block
Sensory Module
Analog Process Module
A/D
Comm.
Module
DSP
Module
Driving
Module
Wireless I/F
BioSignal I/F
Control Module
DSP Module
Comm.
Module
Brain/Body
Wearable BCI
BioData I/F
External Device
Application
Command API
Command API
Command API
Command API
Command API
Event/Drive
External Device
Command API
Command/Response
Event/Drive
Event/Drive
Event/Drive

## Slide 13: Command/Response API
Filter
Turn ON/OFF
Power Level Setting
Signal Threshold Setting
Timing
DSP Function
Data Encoding
Module Type
….

## Slide 14: Functioning Blocks
Sensory Signal Det.
Encoding
Decoding &
Signal Processing
Controlling Ext. Devices
Instrument Amp.
A/D
Wireless Mod.
Wireless Demod.

## Slide 15: Sensory Signal Detect
ECoG
EMG
EEG
Sensing
Pre-Processing
Implantable-Electrode Array Bio-Patch Array Bio-Sensor Array
(EEG, EMG, ECoG)
Contact -Conductance Control
Noise,
CrossTalk,
Low-signal,
Sensing Array

## Slide 16: Pre-Processing and Amplification
Balanced Pre-Amplification
Low-Noise Inst. Amp.
Balanced Differential Amp.
Open-Ground
No Temperature Rise
Pre-Processing
Analog Filter w/ Low-signal
Matching Circuitry
Multi-Channel Input
Multiplexing
Refine Signal Amplified
Noise Reduction

## Slide 17: A/D Conversion
A/D Conversion
Signal-Level Prediction
Sampling/Data Rate
Post-Processing
Digital Control
Digital Data Rate
Encoding
Digital Interface
VCSEL
SPI
External Transmission

## Slide 18: External Control
Controller
Decoding
Signal Processing(DSP)
Pattern Prediction
SmartPhone Application
(Bluetooth)
External Receiving
Wireless Transmission
Blurtooth BLE
ZigBee
Optical Coupling (VCSEL/Photo Cell)
Decision for Analysis/Actuator
Pattern Monitoring
Actuator/ Robotics Control

## Slide 19: External Control
Smart Phone Application
(Wireless Internet)
External Device Control
Robotics
Actuator
Monitoring

## Slide 20: Break-down of considerations Sensory front-end
Sensor type
  Implantable Electrode(ECog)
  External bio-patch(EMG, EEG)
Number of sensors in array for multi-channel processing and multiplexing
Ground type
  Open ground or Common ground
Balanced input or Single input

## Slide 21: Break-down of considerations Sensing & Pre-Processing
Signal volatage level targeted
  10uV ~ 300uV
  Noise level allowed for LNA
Channel separation by filter
Matching circuitry for ohmic contact to skin or brain (contact conductance control)
Cabling type for mechanical connecting
  Flexible cables, Hard wires or Bare Wires
  Connecting Methods

## Slide 22: Break-down of considerations Pre-Amplification
Chip-set for low noise instrument amp.
  Analog Device, Burr Brown
  Intan (Integrated one of DSP)
Balanced input or single input
Ground type
# of channels
Power supplying
Low temperature rise when applying inst.amp.
Multiplexing the amplified outputs
Noise reduction by filter

## Slide 23: Break-down of considerations A/D Conversion
Input signal level prediction 1~100mV
  Powe Budget Analysis
Sampling rate and data rate
Chip set of A/D Convter
A/D Conversion control circuitry
Power supplying
# of channels A/D conversion
Digital control logic

## Slide 24: Break-down of considerations A/D Conversion(cont.)
Digital control logic for transmit control
Digital data rate
Encoding methods of digital output from A/D
Digital interface
  VCSEL : IR/PhotoCell
  SPI

## Slide 25: Break-down of considerations External Transmission
Wireless Transmission
  Bluetooth Low Energy
  ZigBee
  Optical Coupling
    VCSEL/PhotoCell
Power supplying Options
  Wireless charging circuitry + Small Battery
  Wireless power supplying

## Slide 26: Break-down of considerations External Controller
Digital decoding
Digital signal processing
Pattern prediction
Controller applications in PC
Controller applications in SmartPhone
  Directly applied to SmartPhone by Bluetooth
Analysis programming
Monitoring
External Actuator/Robotics control

## Slide 27: Interface Issues
Characterize each interface for implementation
Availability for the tune-up of interface specifications with various implementations by 3rd party
Power budget analysis required at each interface
Low power consumption design required
  Bluetooth L.E.
  VCSEL Array (Vertical Cavity Surface Emitting Laser)

## Slide 28: Considerations for BCI
3 Phases Approach
1st Phase
  Feasibility for applicability to BCI
  Intan off-shelf chip set + Nordic BLE chip set
2nd Phase
  Miniaturizing the size and Form factor
  Power supply issue: Wireless power or Battery

## Slide 29
3rd Phase
  Mount Die-Chips on Flexible PCB
  Integrate analog parts with electrodes in front sensor end
  Integrate digital parts with wireless antenna
  SOC

## Slide 30: 1st Phase Considerations
Intan chip(RHD2232) + Nordic BLE chip(nRF51822) by off-the-shelf products
Miniaturized precision connector or direct cabling linked to bio-patches for collecting bio-signals
Chip-antenna of 2.4GHz for Nordic BLE nRF51822
Polymer lithium battery for power supply with flexibility

## Slide 31: 1st Phase Considerations
Power budget analysis at each stage
Power supply alternative options with wireless power supply by induction coil printed in PCB
Temperature rise by induction coil (~2°K)
Microwave activity influence in brain
Flexible PCB applicability

## Slide 32: Implementation Issues
Implementation Issues
Sensory Front-End
Amp
Pre-Filter
Post-Filter
I/O
A/D
D/A
Wireless RF
Wireless Optic
Wireless RF
Wireless Optic
Signal Processing
Bio-Patch
Intan RHD2232
nRF51822
nRF51822
Smart-Phone
PC
Bluetooth L.E.
Laser Optics
VCSEL
VCSEL
Filter BW Cont., Optic Encoding
Power Supply Circuitry
Polymer-Ion
Battery
Wireless Charging Circuitry
20mm X 15(<20)mm
Sensory Front-End
Sensor Array

## Slide 33: Implementation
SPI
Intan
RHD
2232
nRF
51822
10mm
10mm
Bio-Patch
Analog Part
Digital & Wireless
(Bluetooth L.E.)
30mm
Antenna